## Slide 1: CERN RFD HOM Measurements
On behalf of SY-RF-AC and HL-LHC WP4

## Slide 2: Outline
RFD Crab cavity HOM couplers
Z0 = 25 ohm
25 ohm impedance adapters.
Pre-installation test-box measurements.
HOM measurements of the jacketed cavities at 2 K.

## Slide 3: CERN RFD crab cavities
Crabbing in the horizontal plane.
2 cavities (inc. couplers) manufactured at CERN for tests in the SPS.
- 2 mm
RFD #1
RFD #2
2 mm shorter than nominal design to tune the fundamental mode frequency.

## Slide 4: RFD HOM couplers
V-HOMC
Damps HOMs with high Z⊥v and Z∥
100 Wp range
H-HOMC
Damps HOMs with high Z⊥h and Z∥
kWp range
Field Antenna (FA)
Couples 1 W of the 400 MHz signal.

## Slide 5: RFD HOM couplers
H-HOMC
V-HOMC
FA
… with huge thanks to our colleagues in EN-MME!

## Slide 6: Feedthroughs
25 Ω to increase the diameter of inner conductor…
Verified in simulations to give good HOM damping.
Thermal and mechanical shock tests to arrive at a standard feedthrough design.
With Al2O3, the inner diameter increases by a factor of 3.7 if we move to 25 Ω!
See talk by E. Montesinos for more details.

## Slide 7: Impedance adapters
25 to 50 Ω adapters designed to qualify the HOM couplers and antennas before and after installation onto the cavity.
Developed with...
See talk by K. Turaj for more details.
OD = 40 mm (25 Ω)  N-type (50 Ω)
VHOMC and HHOMC (inc. test box).
Taper for broadband impedance transformation.
~ 450 mm
OD = 16 mm (25 Ω)  N-type (50 Ω)
Field antenna.
Quarter wave transformer (400 MHz)
~ 300 mm
Could also use taper for broadband, but longer.
25 Ω
25 Ω
50 Ω
50 Ω
400 MHz

## Slide 8: Impedance adapters
25 to 50 Ω adapters designed to qualify the HOM couplers and antennas before and after installation onto the cavity.
‘Back-to-back’ measurements qualified the adapters functionality
Developed with...
See talk by K. Turaj for more details.
Back-to-back measurements
Adapter V/H-HOMC:
S21 > -0.5 dB broad-band
Adapter FA:
S21 > -0.1 dB at 400 MHz

## Slide 9: HHOMC Test-box
Text boxes designed to test and quantify the functionality of the couplers.
At frequencies < 1.2 GHz, the curves fit well within the mask: Qualified most important frequency range.
At high frequencies, the couplers are very close to the designed response.
Looking into measurement technique to validate at high frequencies.
The similarity of the three couplers shows very good manufacturing consistency.
HHOMC 002
HHOMC 003
HHOMC 001
Developed with...

## Slide 10: HHOMC Test-box
HHOMC 002
HHOMC 003
HHOMC 001
Developed with...
Fundamental mode rejection sufficient for all three couplers.
Fantastic achievement from our colleagues in EN-MME!
Text boxes designed to test and quantify the functionality of the couplers.
At frequencies < 1.2 GHz, the curves fit well within the mask: Qualified most important frequency range.
At high frequencies, the couplers are very close to the designed response.
Looking into measurement technique to validate at high frequencies.
The similarity of the three couplers shows very good manufacturing consistency.

## Slide 11: Vertical tests
RFD2
Field antenna impedance adapter contact problem.
Non-conformity report. [1]
HOM measurements taken between the HHOMC and VHOMC. Qualified.
Adapter fixed, no second test necessary for cavity qualification.
RFD1
APRIL) Test 1: Leak (ancillaries changed to spare set afterwards)
NO HOM measurements
MAY) Test 2: Second vacuum leak (new ancillaries)
NO HOM measurements
AUG) Test 3: Successfully tested.
HOM measurements taken at 2 K. Qualified.
See talk by K. Turaj for more details.

## Slide 12: Transmission measurements
S21 measurements taken between the ancillaries at 300 K and 2 K – Agrees very well with simulations!
No measurements with Field Antenna for RFD #2 at 2 K due to impedance adapter non-conformity [1].
See talks by E. Montesinos and K. Turaj for more details.

## Slide 13: HOM measurements
f Vs Q for both cavities, modes spectra very similar to simulations.
High frequency deviations expected (cavity tuned for f0).
Acceptance criteria: catalogue mode parameters.

## Slide 14: HOM measurements
High impedance modes tracked for cataloguing and analysis.
Acceptance criteria: Simulation / measurement deviations for high Z modes submitted to beam dynamics experts for further analysis.

## Slide 15: High power HOM
[2]
Depending on the mode parameters, interaction between the 750 MHz longitudinal mode and the 19th bunch spacing harmonic of HL-LHC type beam can produce large heat load.
Strict frequency and quality factor limits placed on this mode.
Both cavities within threshold! HOM power << 1 kW and within specification.
Mode frequency and Q-factor reproducible.
Again, congratulations to SY-RF-AC and EN-MME for both the cavity (mode frequency) and coupler (mode external-Q) manufacture.
No risk of high heat loads for this mode!

## Slide 16: Fundamental mode rejection
Fundamental mode leakage limit of < 6.7 W at VT = 4.1 MV
All meet acceptance criteria!
This shows the validity of the HOM coupler test box as a qualification test.
[2]

## Slide 17: Conclusions
Impedance adapters designed and manufactured.
HHOMC test-box qualified couplers and allowed selection criteria
HOM measurements align with simulations
High power HOM within specification
Fundamental mode rejection within specification
Agree below 1.2 GHz, very consistent manufacture.
High frequency qualification on-going.
Important HOMs tracked and measured.
Maximum HOM power < 1 kW with HL-LHC beam
< 10 W @ VT = 5 MV
In collaboration with and thanks to:
Qualified before installation on the test-box.
Used on the cavity and ‘posed no issue’.

## Slide 18: Thanks! Questions …
In collaboration with and thanks to:
Thanks my management for the great opportunities.
Thanks to Rama, Ofelia and the HiLumi project management for such interesting work.
Thanks to Seb, Antoine, Stefanie, our colleagues in SM18 and in EN-MME for the amazing work, collaboration and fruitful discussions.
H-HOMC
V-HOMC
FA

## Slide 19: References
[1] Non-conformity report for 10-16 adapter: EDMS 1501109
[2] Crab Cavity HOMs - Details and Qualifications: EDMS 2488213